Cable evaluation

Objective:
Evaluate whether a thicker or thinner cable provides better EEG signal quality, comparing electrodes FP1 (thick cable) and FP2 (thin cable) across two conditions: eyes open (EO) and eyes closed (EC).

Protocol:

- Two spiky dry electrodes were positioned simultaneously at FP1 and FP2.

- FP1 used a regular 26 AWG cable, while FP2 used a thinner 28 AWG cable.

- Reference electrodes were connected using 26 AWG cables and spiky electrodes.

- Three recordings were taken for each condition (eyes closed and eyes open), each lasting 40 seconds.

Procedure:

- Impedance Check:
Using OpenBCI, impedances for both FP1 and FP2 electrodes were monitored and found to oscillate around 300 kΩ.

- Recordings:
Three recordings were made for each condition (eyes closed and eyes open), each lasting 30 seconds. The first 3 seconds of each recording were excluded to avoid artifacts caused by the delay in closing the eyes (as the recording was self-administered).

- Data Analysis:
Visual inspection of the recorded signals was performed.

Eyes open

|  | EO1 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| ALPHA POWER | 86.961 uV^2 | 109.841 uV^2 |
| BETA POWER |  |  |
| NOISE |  |  |
| RELATIVE alpha POWER | 0.072 | 0.067 |
| NOISE RATIO | {'delta': 1.2328245966276967, 'theta': 0.8598885517348224, 'alpha': 0.6612346807533896, 'beta': 1.8216936997841868} | {'delta': 2.622152774538213, 'theta': 1.3529088958245938, 'alpha': 0.8798675944354596, 'beta': 2.3634700889876408} |
| METRICS QUALITY | {'alpha_power_uV2': 86.96074239242836, 'beta_power_uV2': 251.4803766878843, 'delta_power_uV2': 159.03375002743027, 'line_noise_power_uV2': 98.41613038730944, 'zero_crossings_per_sec': 57.304038004750595, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 186.6911782781535, 'symmetry_corr_LR': 0.8600784682619796} | {'alpha_power_uV2': 109.84094286713527, 'beta_power_uV2': 313.1901092077789, 'delta_power_uV2': 325.909877668459, 'line_noise_power_uV2': 125.59475691781392, 'zero_crossings_per_sec': 55.819477434679335, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 221.74598134864314, 'symmetry_corr_LR': 0.8600784682619798} |

Eyes closed

|  | EO2 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| ALPHA POWER |  |  |
| BETA POWER |  |  |
| NOISE |  |  |
| RELATIVE alpha POWER |  |  |
| NOISE RATIO | {'delta': 0.9755229093069282, 'theta': 1.0445293255230315, 'alpha': 0.7225529847411846, 'beta': 2.2546642631652194} | {'delta': 2.1323337702862752, 'theta': 1.4433429983197053, 'alpha': 0.9078763268222189, 'beta': 2.6826973965254663} |
| METRICS QUALITY | {'alpha_power_uV2': 76.20300845598565, 'beta_power_uV2': 255.4812725692296, 'delta_power_uV2': 95.20698560466417, 'line_noise_power_uV2': 63.21071861820434, 'zero_crossings_per_sec': 57.2700296735905, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 173.91972086546912, 'symmetry_corr_LR': 0.8775689615018331} | {'alpha_power_uV2': 97.38177457590787, 'beta_power_uV2': 308.5681574593464, 'delta_power_uV2': 219.20671502094166, 'line_noise_power_uV2': 100.91696708374562, 'zero_crossings_per_sec': 55.97181008902077, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 222.1873004822329, 'symmetry_corr_LR': 0.8775689615018331} |

|  | EO3 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| NOISE |  |  |
| NOISE RATIO | {'delta': 0.9022061179241994, 'theta': 0.9891695769814548, 'alpha': 0.7071475076064468, 'beta': 2.5052144308651427} | {'delta': 1.6172750650401952, 'theta': 1.3114675087579606, 'alpha': 0.8696980148178305, 'beta': 3.096810385456438} |
| METRICS QUALITY | {'alpha_power_uV2': 80.47412008808158, 'beta_power_uV2': 303.0582715669054, 'delta_power_uV2': 97.69934887542202, 'line_noise_power_uV2': 42.983107951834, 'zero_crossings_per_sec': 58.25783454626467, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 157.16706163822448, 'symmetry_corr_LR': 0.9169353976670045} | {'alpha_power_uV2': 98.66242641578874, 'beta_power_uV2': 374.5609506291501, 'delta_power_uV2': 178.23782231494306, 'line_noise_power_uV2': 66.07208486328332, 'zero_crossings_per_sec': 55.80721817911778, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 196.14216869813603, 'symmetry_corr_LR': 0.9169353976670046} |

|  | EC1 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| NOISE |  |  |
| NOISE RATIO | {'delta': 1.2005789726772027, 'theta': 1.2655556744778524, 'alpha': 1.063702792496175, 'beta': 2.9267036803099953} | {'delta': 1.9571361684365198, 'theta': 1.527146209811782, 'alpha': 1.2000139533678886, 'beta': 3.3332595993399283} |
| METRICS QUALITY | {'alpha_power_uV2': 107.5961831206613, 'beta_power_uV2': 295.40228378702375, 'delta_power_uV2': 112.41682300389607, 'line_noise_power_uV2': 14.97341099017701, 'zero_crossings_per_sec': 51.135520261243876, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 141.0115860049683, 'symmetry_corr_LR': 0.9589090100175088} | {'alpha_power_uV2': 128.41681024660696, 'beta_power_uV2': 358.7383074737281, 'delta_power_uV2': 193.3645013202345, 'line_noise_power_uV2': 21.062401874174537, 'zero_crossings_per_sec': 51.469496808668545, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 175.04940411149335, 'symmetry_corr_LR': 0.9589090100175087} |

|  | EC2 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| NOISE |  |  |
| NOISE RATIO | {'delta': 0.6887075424031442, 'theta': 1.290931850800313, 'alpha': 0.785285466242677, 'beta': 2.5315296293589356} | {'delta': 1.0717842707778336, 'theta': 1.8926680257658834, 'alpha': 1.0035650609417404, 'beta': 3.2614230732044276} |
| METRICS QUALITY | {'alpha_power_uV2': 84.68336771389123, 'beta_power_uV2': 285.14334810525963, 'delta_power_uV2': 71.38482561400555, 'line_noise_power_uV2': 50.404782850827644, 'zero_crossings_per_sec': 58.04631828978623, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 147.6947726823022, 'symmetry_corr_LR': 0.9597943922023056} | {'alpha_power_uV2': 110.82971963415841, 'beta_power_uV2': 372.7359040748207, 'delta_power_uV2': 116.89535059987544, 'line_noise_power_uV2': 87.66690944792379, 'zero_crossings_per_sec': 55.59679334916865, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 159.21557383163275, 'symmetry_corr_LR': 0.9597943922023056} |

|  | EC3 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| NOISE |  |  |
| NOISE RATIO | {'delta': 1.2667547209045036, 'theta': 1.3094507247993203, 'alpha': 0.9788340338825489, 'beta': 2.5567444014971006} | {'delta': 2.4685051995057354, 'theta': 1.7867484076253184, 'alpha': 1.1083464895242112, 'beta': 2.8366288539545965} |
| METRICS QUALITY | {'alpha_power_uV2': 93.95787424018539, 'beta_power_uV2': 268.51191181603065, 'delta_power_uV2': 115.13993315571156, 'line_noise_power_uV2': 6.671694106134741, 'zero_crossings_per_sec': 51.58199643493761, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 182.96795673437083, 'symmetry_corr_LR': 0.9085901904927783} | {'alpha_power_uV2': 117.5123943474753, 'beta_power_uV2': 325.19855664707524, 'delta_power_uV2': 245.5173174857619, 'line_noise_power_uV2': 14.276031671816423, 'zero_crossings_per_sec': 47.27421271538918, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 191.35716972592047, 'symmetry_corr_LR': 0.9085901904927784} |

MEAN and IQR of 3 recordings, ratio band noise

RATIO Eyes open

FP1- thick wire

delta    1.0369 ± 0.1653

theta    0.9645 ± 0.0923

alpha    0.6970 ± 0.0307

beta     2.1939 ± 0.3418

FP2-thin wire

delta    2.1239 ± 0.5024

theta    1.3692 ± 0.0659

alpha    0.8858 ± 0.0191

beta     2.7143 ± 0.3667

Ratio eyes closed

Fp1

delta    1.0520 ± 0.2890

theta    1.2886 ± 0.0219

alpha    0.9426 ± 0.1392

beta     2.6717 ± 0.1976

Fp2

delta    1.8325 ± 0.6984

theta    1.7355 ± 0.1828

alpha    1.1040 ± 0.0982

beta     3.1438 ± 0.2483

Here the ratios show that the relation band/noise is better for the thin cable, which could mean higher attenuation of the noise or best resistance to it. The IQR shows that the recordings are consistent, with a low variability among them an a good reproducibility among the 3 recordings.

METRICS

AVERAGE OF 3 recordings

| EOFP1 | EOFP2 |
| --- | --- |
| alpha_power_uV2 81.21 beta_power_uV2 270.01 delta_power_uV2 117.31 line_noise_power_uV2 68.20 zero_crossings_per_sec 57.61 percentage_flatline 0.00 large_jumps 0.00 range_uV 172.59 symmetry_corr_LR 0.88 dtype: float64 | alpha_power_uV2 101.96 beta_power_uV2 332.11 delta_power_uV2 241.12 line_noise_power_uV2 97.53 zero_crossings_per_sec 55.87 percentage_flatline 0.00 large_jumps 0.00 range_uV 213.36 symmetry_corr_LR 0.88 dtype: float64 |
| ECFP1 | ECFP2 |
| alpha_power_uV2 95.41 beta_power_uV2 283.02 delta_power_uV2 99.65 line_noise_power_uV2 24.02 zero_crossings_per_sec 53.59 percentage_flatline 0.00 large_jumps 0.00 range_uV 157.22 symmetry_corr_LR 0.94 | alpha_power_uV2 118.92 beta_power_uV2 352.22 delta_power_uV2 185.26 line_noise_power_uV2 41.00 zero_crossings_per_sec 51.45 percentage_flatline 0.00 large_jumps 0.00 range_uV 175.21 symmetry_corr_LR 0.94 |

These results suggest that the captured signal is better for the thin cable, which align with the signal to noise ratio computations.

The noise power is higher for both conditions, which could suggest that the captured noise is higher. However, together with the computed noise ratios, the util signal band also grows, being favourable for the thin cable.

Zero crossing are quite similar for both cables

Symmetry correlation shows that the registered signal is consistent for both sides

STATS

EYES CLOSED

With eyes closed, alpha and beta have significant differences, with higher values in FP2. WE only work with 3 recordings so we can only expect a trend

EYES OPEN

With eyes open, we also observe a trend with higher values for fp1 than fp2

STAT DESCRIPTIVE

EYES CLOSED

EYES OPEN

Key Findings:

- Signal-to-Noise Ratios (Ratio of Band Power to Noise > 55 Hz):

  - Across all frequency bands (delta, theta, alpha, beta) and conditions, the thin cable (FP2) consistently showed higher signal-to-noise ratios than the thick cable (FP1).

  - This suggests that the thin cable transmits EEG signals with better relative quality compared to the thick cable.

- Data Variability (IQR):

  - The interquartile range (IQR) values were generally small to moderate, indicating consistent and reproducible results across the 3 recordings.

  - Some variability exists but is not enough to undermine the observed trend favoring the thin cable.

- Power Metrics (Average µV²):

  - The thin cable yielded higher absolute power in alpha, beta, and delta bands across both EO and EC conditions, indicating a stronger EEG signal.

  - Although line noise power was also higher with the thin cable, the increase in signal power was proportionally greater, resulting in an overall improved signal-to-noise scenario.

- Signal Quality Indicators:

  - Metrics like zero crossings per second, percentage of flatlines, and large jumps showed no significant difference between cables, and very low artifact presence, supporting the reliability of both setups.

- Amplitude Range:

  - The thin cable recordings had a wider amplitude range, further supporting a stronger EEG signal capture.

- Symmetry Correlation Between Hemispheres:

  - High correlation values (~0.88-0.94) between FP1 and FP2 electrodes suggest highly consistent and symmetric brain activity measurements.

  - This confirms the reliability of recordings from both cables with no apparent degradation in signal coherence on either side.

Other tests that could be perfomed: Improvements

- Coherence: computing coherence with nearby electrodes, however we don’t have enough electrodes to perform this analysis

- Phase lag: to check if cable introduce delays

- Signal to artifact

Good practices

- Stats, more recordings, more subjects

- Cable mechanical stress, important for the design and mechanical part

- Environmental noise susceptibility ( near power lines, indoors and outdoors, checking shielding)

- Longer duration recordings