Cable evaluation

Protocol

- Two spiky dry electrodes were positions in fp1 and fp2 at the same time

- Fp1 using the regular cable 26 awg and fp2 using cable 28 awg

- Reference electrodes were using the 26 awg cable and spi ky electrodes

- 3 recordings eyes closed, 3 recordings eyes open, 40 seconds

- Checking impedances

  - Openbci checking: both impendaces fp1 and fp2 oscillate around 300 kohm

- 3 recordings eyes closed and eyes open, 30 seconds, 3 first seconds could be cut in orer to avoud artifacts of the delay in closing the eyes ( as I recorded myself)

- Visual inspection

Eyes open

|  | EO1 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| ALPHA POWER | 86.961 uV^2 | 109.841 uV^2 |
| BETA POWER |  |  |
| NOISE |  |  |
| RELATIVE alpha POWER | 0.072 | 0.067 |
| NOISE RATIO | {'delta': 1.2328245966276967, 'theta': 0.8598885517348224, 'alpha': 0.6612346807533896, 'beta': 1.8216936997841868} | {'delta': 2.622152774538213, 'theta': 1.3529088958245938, 'alpha': 0.8798675944354596, 'beta': 2.3634700889876408} |
| METRICS QUALITY | {'alpha_power_uV2': 86.96074239242836, 'beta_power_uV2': 251.4803766878843, 'delta_power_uV2': 159.03375002743027, 'line_noise_power_uV2': 98.41613038730944, 'zero_crossings_per_sec': 57.304038004750595, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 186.6911782781535, 'symmetry_corr_LR': 0.8600784682619796} | {'alpha_power_uV2': 109.84094286713527, 'beta_power_uV2': 313.1901092077789, 'delta_power_uV2': 325.909877668459, 'line_noise_power_uV2': 125.59475691781392, 'zero_crossings_per_sec': 55.819477434679335, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 221.74598134864314, 'symmetry_corr_LR': 0.8600784682619798} |

Eyes closed

|  | EO2 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| ALPHA POWER |  |  |
| BETA POWER |  |  |
| NOISE |  |  |
| RELATIVE alpha POWER |  |  |
| NOISE RATIO | {'delta': 0.9755229093069282, 'theta': 1.0445293255230315, 'alpha': 0.7225529847411846, 'beta': 2.2546642631652194} | {'delta': 2.1323337702862752, 'theta': 1.4433429983197053, 'alpha': 0.9078763268222189, 'beta': 2.6826973965254663} |
| METRICS QUALITY | {'alpha_power_uV2': 76.20300845598565, 'beta_power_uV2': 255.4812725692296, 'delta_power_uV2': 95.20698560466417, 'line_noise_power_uV2': 63.21071861820434, 'zero_crossings_per_sec': 57.2700296735905, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 173.91972086546912, 'symmetry_corr_LR': 0.8775689615018331} | {'alpha_power_uV2': 97.38177457590787, 'beta_power_uV2': 308.5681574593464, 'delta_power_uV2': 219.20671502094166, 'line_noise_power_uV2': 100.91696708374562, 'zero_crossings_per_sec': 55.97181008902077, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 222.1873004822329, 'symmetry_corr_LR': 0.8775689615018331} |

|  | EO3 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| NOISE |  |  |
| NOISE RATIO | {'delta': 0.9022061179241994, 'theta': 0.9891695769814548, 'alpha': 0.7071475076064468, 'beta': 2.5052144308651427} | {'delta': 1.6172750650401952, 'theta': 1.3114675087579606, 'alpha': 0.8696980148178305, 'beta': 3.096810385456438} |
| METRICS QUALITY | {'alpha_power_uV2': 80.47412008808158, 'beta_power_uV2': 303.0582715669054, 'delta_power_uV2': 97.69934887542202, 'line_noise_power_uV2': 42.983107951834, 'zero_crossings_per_sec': 58.25783454626467, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 157.16706163822448, 'symmetry_corr_LR': 0.9169353976670045} | {'alpha_power_uV2': 98.66242641578874, 'beta_power_uV2': 374.5609506291501, 'delta_power_uV2': 178.23782231494306, 'line_noise_power_uV2': 66.07208486328332, 'zero_crossings_per_sec': 55.80721817911778, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 196.14216869813603, 'symmetry_corr_LR': 0.9169353976670046} |

|  | EC1 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| NOISE |  |  |
| NOISE RATIO | {'delta': 1.2005789726772027, 'theta': 1.2655556744778524, 'alpha': 1.063702792496175, 'beta': 2.9267036803099953} | {'delta': 1.9571361684365198, 'theta': 1.527146209811782, 'alpha': 1.2000139533678886, 'beta': 3.3332595993399283} |
| METRICS QUALITY | {'alpha_power_uV2': 107.5961831206613, 'beta_power_uV2': 295.40228378702375, 'delta_power_uV2': 112.41682300389607, 'line_noise_power_uV2': 14.97341099017701, 'zero_crossings_per_sec': 51.135520261243876, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 141.0115860049683, 'symmetry_corr_LR': 0.9589090100175088} | {'alpha_power_uV2': 128.41681024660696, 'beta_power_uV2': 358.7383074737281, 'delta_power_uV2': 193.3645013202345, 'line_noise_power_uV2': 21.062401874174537, 'zero_crossings_per_sec': 51.469496808668545, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 175.04940411149335, 'symmetry_corr_LR': 0.9589090100175087} |

|  | EC2 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| NOISE |  |  |
| NOISE RATIO | {'delta': 0.6887075424031442, 'theta': 1.290931850800313, 'alpha': 0.785285466242677, 'beta': 2.5315296293589356} | {'delta': 1.0717842707778336, 'theta': 1.8926680257658834, 'alpha': 1.0035650609417404, 'beta': 3.2614230732044276} |
| METRICS QUALITY | {'alpha_power_uV2': 84.68336771389123, 'beta_power_uV2': 285.14334810525963, 'delta_power_uV2': 71.38482561400555, 'line_noise_power_uV2': 50.404782850827644, 'zero_crossings_per_sec': 58.04631828978623, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 147.6947726823022, 'symmetry_corr_LR': 0.9597943922023056} | {'alpha_power_uV2': 110.82971963415841, 'beta_power_uV2': 372.7359040748207, 'delta_power_uV2': 116.89535059987544, 'line_noise_power_uV2': 87.66690944792379, 'zero_crossings_per_sec': 55.59679334916865, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 159.21557383163275, 'symmetry_corr_LR': 0.9597943922023056} |

|  | EC3 |  |
| --- | --- | --- |
|  | FP1 | Fp2 |
| PSD |  |  |
| NOISE |  |  |
| NOISE RATIO | {'delta': 1.2667547209045036, 'theta': 1.3094507247993203, 'alpha': 0.9788340338825489, 'beta': 2.5567444014971006} | {'delta': 2.4685051995057354, 'theta': 1.7867484076253184, 'alpha': 1.1083464895242112, 'beta': 2.8366288539545965} |
| METRICS QUALITY | {'alpha_power_uV2': 93.95787424018539, 'beta_power_uV2': 268.51191181603065, 'delta_power_uV2': 115.13993315571156, 'line_noise_power_uV2': 6.671694106134741, 'zero_crossings_per_sec': 51.58199643493761, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 182.96795673437083, 'symmetry_corr_LR': 0.9085901904927783} | {'alpha_power_uV2': 117.5123943474753, 'beta_power_uV2': 325.19855664707524, 'delta_power_uV2': 245.5173174857619, 'line_noise_power_uV2': 14.276031671816423, 'zero_crossings_per_sec': 47.27421271538918, 'percentage_flatline': 0.0, 'large_jumps': 0, 'range_uV': 191.35716972592047, 'symmetry_corr_LR': 0.9085901904927784} |

MEAN and IQR of 3 recordings

RATIO Eyes open

FP1- thick wire

delta    1.0369 ± 0.1653

theta    0.9645 ± 0.0923

alpha    0.6970 ± 0.0307

beta     2.1939 ± 0.3418

FP2-thin wire

delta    2.1239 ± 0.5024

theta    1.3692 ± 0.0659

alpha    0.8858 ± 0.0191

beta     2.7143 ± 0.3667

Ratio eyes closed

Fp1

delta    1.0520 ± 0.2890

theta    1.2886 ± 0.0219

alpha    0.9426 ± 0.1392

beta     2.6717 ± 0.1976

Fp2

delta    1.8325 ± 0.6984

theta    1.7355 ± 0.1828

alpha    1.1040 ± 0.0982

beta     3.1438 ± 0.2483

METRICS

AVERAGE OF 3 recordings

| EOFP1 | EOFP2 |
| --- | --- |
| alpha_power_uV2 81.21 beta_power_uV2 270.01 delta_power_uV2 117.31 line_noise_power_uV2 68.20 zero_crossings_per_sec 57.61 percentage_flatline 0.00 large_jumps 0.00 range_uV 172.59 symmetry_corr_LR 0.88 dtype: float64 | alpha_power_uV2 101.96 beta_power_uV2 332.11 delta_power_uV2 241.12 line_noise_power_uV2 97.53 zero_crossings_per_sec 55.87 percentage_flatline 0.00 large_jumps 0.00 range_uV 213.36 symmetry_corr_LR 0.88 dtype: float64 |
| ECFP1 | ECFP2 |
| alpha_power_uV2 95.41 beta_power_uV2 283.02 delta_power_uV2 99.65 line_noise_power_uV2 24.02 zero_crossings_per_sec 53.59 percentage_flatline 0.00 large_jumps 0.00 range_uV 157.22 symmetry_corr_LR 0.94 | alpha_power_uV2 118.92 beta_power_uV2 352.22 delta_power_uV2 185.26 line_noise_power_uV2 41.00 zero_crossings_per_sec 51.45 percentage_flatline 0.00 large_jumps 0.00 range_uV 175.21 symmetry_corr_LR 0.94 |

STATS

EYES CLOSED

With eyes closed, alpha and beta have significant differences, with higher values in FP1. WE only work with 3 recordings so we can only expect a trend

EYES OPEN

With eyes open, we also observe a trend with higher values for fp1 than fp2

STAT DESCRIPTIVE

EYES CLOSED

EYES OPEN